## Slide 1
Measurement Systems Analysis
Components and Acceptance Criteria

## Slide 2: Purpose
To understand key concepts of measurement systems analysis
To understand potential sources of measurement error and its affect on both process and part quality

## Slide 3: Why do we care about Measurement Systems Analysis(MSA)?
Ensures measurement is accurate and precise
Eliminates false data that can impact decisions about the process
Provides information about the gage(s) being used to evaluate our parts and processes
Reduces risk in calling good parts bad (Type I Error) **Impacts internal throughput, scrap and rework
Reduces risk in calling bad parts good (Type II Error) **Impacts customer
Data you can trust = Good Decisions

## Slide 4: Which line has less variation
Chart: 
Chart: 
Line 1
Line 2

## Slide 5: Which line has less variation?
Chart: 
Chart: 
Line 1
Line 2
Chart: 
Chart: 
Gage Variation
Total Variation

## Slide 6: Which line has less variation?
Chart: 
Chart: 
Line 1
Chart: 
Chart: 
Gage Variation
Total Variation
Chart: 
Chart: 
Part Variation
Part Variation
+
Gage Variation
=
Total Variation
(Observed)
Line 2 has
Less variation
Line 2

## Slide 7: Understanding Variation
Chart: 
Chart: 
Chart: 
Part Variation
+
Gage Variation
=
Total Variation
(Observed)
What is the best course of action?
Assume the total observed variation is below
1.33 Ppk for an initial process study on a starred
dimension (1.67 Ppk or greater required)
Measure a different sample of parts (Gather more data)
Investigate process changes to reduce process variation
Investigate gage changes to reduce measurement variation

## Slide 8: How do you decide what to do?
How do you decide what to do?
What are typical reactions to measuring a part out of tolerance?
  Adjust processes / implement additional detection methods
  Quarantine / Sort / Scrap parts
  Review PFMEAs and Control Plans
Before doing all this verify the measurement system

## Slide 9: MSA Components
What does a measurement system analysis consist of?

## Slide 10: MSA Components
Resolution (Discrimination)
Accuracy
Linearity
Stablity
Repeatability (Test / Retest)
Reproducibility
Gage R&R

## Slide 11: Sources of Variation
Location of the Measurements
  Accuracy / Bias
  Stability
  Linearity
Variation of Measurement
  Repeatability
  Reproducibility
Resolution / Discrimination
Each source of variation can result in incorrect measurements

## Slide 12: Resolution
Can the gage detect change?
Which logo below is larger?
How is the measurement different?

## Slide 13: Resolution - Response
Select gages that have greater levels of discrimination (smaller increments)
Record all measurement data, i.e. if a gage measures to 3 decimal points record all three
As a general rule gage discrimination should be 1/10th of the product tolerance
  Product tolerance: 0.100
  Gage tolerance: 0.010 (Recommended minimum)

## Slide 14: Accuracy
What is accuracy?
The difference between the measured average value and the master value or the ‘Shift’ from ‘True’ value
The master value needs to come from a traceable measurement source (ex. CMM)
Although accuracy and bias are often used interchangeably it is best practice to refer to this difference in values as bias

## Slide 15: Accuracy
What is accuracy?
Measured values
Master value
(reference standard)
Let’s assume the master value is 5
Let’s also assume the average measurement is 5.5
In this case the accuracy of the gage is ‘bias’ by 0.5 above the master

## Slide 16: How to Determine Bias
Take three samples
  Ideally these samples would be at the high, middle and low ends of the process
Determine the reference value for each sample using a traceable measurement method (CMM, calibrated height gage, etc.)
Measure each sample at least 10 times. Calculate the average and subtract from the reference value
Determine if the accuracy is acceptable or needs to be corrected
  This is to be determined by the working team

## Slide 17: Bias Example
Note: With multiple samples a linearity study may be completed

## Slide 18: Potential Bias Causes
Gage needs calibration
Worn gage, equipment or fixture
Error in master part (worn or damaged)
Improper setting of master in gage
Differing measurement methods (set, loading, clamping, technique, etc.)
Environment (temperature, humidity, vibration, cleanliness, etc.)

## Slide 19: Bias - Response
Increase calibration schedule
Implement gage operating instructions
Verify gage was made to correct specifications
Check gage for wear
Cleaning of gage
Evaluate gage environment

## Slide 20: Linearity
What is linearity?
  Linearity is the change in bias across the operating range of the gage

## Slide 21: Linearity
Chart: 
 | 1 | 2.4 | 3.8 | 5.2 | 6.6 | 8 | 9.4 | 10.8 | 12.2 | 13.6
 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Bias
No Bias
(Master value equals measured average value)
(Difference between the measured average value and the master value)
Full Range of Gage (Master Value)

## Slide 22: How to Determine Linearity
Pick 5 parts covering the entire product tolerance
Determine reference value of each part (ex. CMM)
Measure each part 10 times or more and determine averages
Subtract the average values from the determined reference values and decide if linearity is acceptable

## Slide 23: Linearity Calculation (Minitab)
1. Enter Part Number, Measurement and Reference Value in the Worksheet
2. Go to ‘Stat’, ‘Quality Tools’, ‘Gage Study’, ‘Gage Linearity and Bias Study’
Fill in column ID’s for Part Number, Reference Values and Measurement Data. Don’t forget to fill in the Gage Info!

## Slide 24: Linearity Analysis (Minitab)
This gage has little bias between 1.0 and 1.2
However, there is significant bias (P<0.05) at 1.31 and 1.35

## Slide 25: Linearity – Acceptance Criteria
Study must use 5 or more parts
Parts must cover entire operating range of gage
A p-value of less than 0.05 (P<0.05) indicates statistically significant bias

## Slide 26: Linearity - Response
Implement correction factor / table / curve
Review gage and adjust
Restrict use of range (Need to consider specification limits)
Verify proper calibration (through entire operating range)
Verify part / fixture is not distorted with change in part size

## Slide 27: Linearity Question
Question: Is this linearity study acceptable?
Yes, all points are statistically acceptable across the range
No, at 1.355 there is significant bias
No, at 1.315 and 1.355 there is significant bias
There is no bias, this is a linearity study

## Slide 28: Stability
What is stability?
  The variation in measurements obtained with the same master part(s) on a single feature over a period of time
  Also known as ‘measurement drift’ and the ‘change in bias over time’

## Slide 29: Stability
Measurements taken Monday
Measurements taken Friday
Measurement results are changing over time. Gage is not stable

## Slide 30: How to Determine Stability
Select a master part
  Part may be a reference part (master) or a production part
  You can select multiple parts (low, middle and high end of expected range) and track stability for each
Determine the reference value for each part via a master measurement method
Measure each part 3 – 5 times at a chosen interval (daily, weekly, etc.)
  Frequency should be determined based on knowledge of the gage. If frequency is unclear you should start at a high frequency and reduce over time as deemed acceptable.
Plot the ongoing data on a Xbar & R chart

## Slide 31: Stability – Potential Causes
Calibration issues
Gage wear
Poor maintenance
Master part damage / wear

## Slide 32: Stability - Response
Increase frequency of calibration
Increase frequency of PM
Change / Adjust gage
Establish ‘life’ of gage or master part

## Slide 33: Stability Question
Question: Assuming the same part was measured is this gage stable?
Can this gage be used?
Process tolerance is 9.0/6.0
Yes, the process is stable. Use the gage.
No, the gage is not stable. Do not use the gage.
No, the gage is not stable. Develop calibration schedule and use the gage.
The gage is not stable.
Larger ‘jumps’ appear to happen every other day. The team should decide how often calibration should take place.

## Slide 34: Precision
What is precision?
  According to AIAG it is “the net effect of discrimination, sensitivity and repeatability over the operating range of the measurement system.” (AIAG MSA Third Edition)
So what does that mean?
  Precision quantifies how close repeated measurements are to each other
Not Precise
Precise

## Slide 35: Precision - Components
Repeatability
  Best known as within user variation
  Also encompasses all ‘within’ variation
    Within part
    Within gage
    Within standard
    Within method
    Within user
    Within environment
Reproducibility
  Best known as between user variation
  Also encompasses all ‘between’ variation
    Between systems
    Between conditions
    Between methods
    Between Environment

## Slide 36: Precision - Components
Chart: 
Chart: 
Operator 1
Operator 2
Operator 3
Repeatability
(Within Variation)
Reproducibility
(Between Variation)

## Slide 37: Variable Data GRR

## Slide 38: Variable Data Gage R&R
How do you conduct a Variable Gage R&R?
Determine the need for the study. What do you want to accomplish?
  %Study – For process control
  %Tolerance – For product control
Determine specifics about the study (number of appraisers, number of trials, etc.)
Select appraisers
  Appraisers should be selected that will used the gage in production
Select the sample parts to be used
  Based upon chosen need from study
  If for process control parts must be chosen from process range

## Slide 39: Variable Data Gage R&R
Verify gage resolution is acceptable
Measure parts in random order to prevent ‘measurement memory’
The standard measurement process should be observed by all users (gage instructions)
The study should be observed by the engineer to ensure reliability of the data

## Slide 40: Variable Data Gage R&R
General Rules
  Three appraisers
  Ten parts
  Three measurements per part per appraiser (Ninety total measurements)
  Parts should make use of the entire range of process variation
    This does not mean the entire product tolerance
    A capability study should accompany a completed GRR
  Parts should be collected over time
    Parts collected in a row will usually have low between part variation
  Use ANOVA method if possible
    Xbar & R method does not account for operator interaction
  Report %Study or %Tolerance (whichever applies)

## Slide 41: Variable Data Gage R&R
General Rules
  Three appraisers
  Ten parts
  Three measurements per part per appraiser (Ninety total measurements)
  Parts should make use of the entire range of process variation
    This does not mean the entire product tolerance
    A capability study should accompany a completed GRR
  Parts should be collected over time
    Parts collected in a row will usually have low between part variation
  Use ANOVA method if possible
    Xbar & R method does not account for operator interaction
  Report %Study or %Tolerance (whichever applies)

## Slide 42: Variable GRR Calculation (Minitab)
1. Enter Appraiser, Part number and Measurement in the Worksheet
2. Go to ‘Stat’, ‘Quality Tools’, ‘Gage Study’, ‘Gage R&R Study (Crossed)’
“Crossed” is when the same parts are used between trials
“Nested” is when different parts are used between trials

## Slide 43: Variable GRR Calculation (Minitab)
Fill in part number, Operators and Measurement Data and select ANOVA
Click Options
Complete Gage Info

## Slide 44: Variable GRR Calculation (Minitab)
Fill in upper and lower limits
and include a title of the GRR
Click OK

## Slide 45: Gage R&R (ANOVA) for IDC

## Slide 46: Variable GRR - Analysis
Pay attention to the overall sample range to understand ‘scale’ of error
Inside the Xbar r Chart control limits is the blind spot of the gage

## Slide 47: Variable GRR - Analysis
Part to part variation should be the largest contributor
R Chart should have all points in control. An out of control condition indicates poor repeatability
Xbar r Chart should be 50% out of control. This indicates each operator can tell a good part from a bad part
Want:
Small spread around mean
Similar mean and spread by operator
Lines to be close to being on top of each other

## Slide 48: Variable GRR – Graphical Metrics
Components of Variation
  Indicates repeatability / reproducibility / parts
  Want low Gage R&R, High part to part variation
R Chart
  Indicates repeatability / resolution
  No outliers are permitted, helps to I.D. unusual measurements
  Plateaus or steps indicate resolution limitation
Xbar r Chart
  Reproducibility / Sensitivity
  Want to see similar patterns for each operator

## Slide 49: Variable GRR – Numerical Metrics
Report number of distinct categories
Report %Study or %Tolerance (Whichever applies) More on this later…

## Slide 50: Variable GRR – Numerical Metrics
%Study or %Tolerance should be below 10%
  Between 10% and 30% may be acceptable
  Above 30% is not acceptable
Number of distinct categories must be at least 5

## Slide 51: Why ANOVA?

## Slide 52: Xbar/R GRR

## Slide 53: ANOVA GRR

## Slide 54: Xbar/R & ANOVA Comparison
The same data was used for both studies
The only difference is method used to analyze the gage
WHY?
The Xbar/R method ignores operator to part interaction and as a result can appear to ‘boost’ gage performance

## Slide 55: Xbar/R & ANOVA Comparison
Xbar/R Method ANOVA Method
Operator interaction is ignored using the Xbar/R method

## Slide 56: Understanding % Tolerance
%Tolerance is the measurement error as a percent of the product specification (tolerance)
%Tolerance determines if the gage can be used for product control (determining a good part from a bad)
If the gage cannot pass %Tolerance it cannot determine if a part within specification

## Slide 57: Understanding %Study
%Study is the measurement error as a percent of total variation (standard deviation)
%Study determines if the gage can be used for process control
If the gage cannot pass %Study it cannot distinguish one part from another within normal process variation or monitor process improvements / process changes

## Slide 58: Specification vs. Total Variation
Product Specification
Total Variation
Chart: 
Chart: 

## Slide 59: Outlier Effect on %Study
Adding an outlier will increase total variation
This will reduce the reported %Study and can make the gage look acceptable when it really isn’t
Chart: 
Product Tolerance

## Slide 60: Outlier Effect on %Tolerance
Adding an outlier will not affect %Tolerance
%Tolerance relies upon the USL-LSL in its calculation. An outlier does not affect the product specification
Chart: 
observed (Total Observed Variation)
Product Tolerance

## Slide 61: Minimizing %Study by Part Selection
%Study is heavily dependent on the parts selected for the study
The more variation between parts selected for the study the lower %Study will be
When selecting parts be sure to include parts across the entire process variation
  Do not select parts outside of the process variation as this can mislead you to think the gage is better than it is

## Slide 62: Part Selection for GRR
%Tolerance is not affected by part selection
%Study is a measure of how well the gage sees inside the range of parts selected. Part selection is critical

## Slide 63: Confidence Interval and Distinct Categories
Confidence Interval
Measurement
Actual Size
The actual size of a measured part can be anywhere within the confidence interval

## Slide 64: Distinct Categories
Distinct categories is the number of confidence intervals (divisions) that the measurement system can accurately measure across the process variation.
Also known as effective resolution (not gage resolution)

## Slide 65: Distinct Categories
Chart: 
Chart: 
Chart: 
Number of distinct categories is an indication of how well a measurement system can detect product variation
Used to establish if the measurement system has acceptable resolution (discrimination)
1 Distinct Category
2-4 Distinct Categories
5+ Distinct Categories

## Slide 66: Confidence Interval and Distinct Categories
Chart: 
Confidence Interval surrounding a measurement
Looking at the measurements what is your estimate of the number of distinct categories?
Conceptually the number of distinct categories is equal to the number of confidence intervals that span the distribution
Number of Distinct Categories is roughly four
Notes: High %Study we can say that we have a wide confidence interval. This results in more overlap of confidence intervals from one measurement to the next.
Low %Study results in narrow confidence intervals, thus we can be more confident that our measurement if giving us good information.

## Slide 67: When to Report % Tolerance
There are cases when you want to report %Tolerance instead of %Study
  When the manufacturing process is stable, in control, and…
  When the parts selected for the GRR cover the entire range of process variation, and…
  When the distribution is not close to the specification limits, and…
  When the process is highly capable (Ppk>=2.5)
  When the dimension is not a SC/CC/ and severity <8

## Slide 68: What Does That Look Like?

## Slide 69: Troubleshooting Variable GRR
Repeatability Issues
  Verify / Review gage instructions
  Verify clamping sequence
  Verify part is not loose in gage
  Verify measuring location is the same between trials
  Verify sufficient gage resolution
  Verify maintenance schedule is sufficient
  Verify operator method does not vary
  Review environmental impact (heat, vibration, etc.)

## Slide 70: Troubleshooting Variable GRR
Reproducibility Issues
  Verify / Review gage instructions
  Verify clamping sequence is the same
  Verify appraiser to appraiser methods are the same
  Appraiser ergonomics (size, strength, etc.)
  Review environmental impact (heat, vibration, etc.)

## Slide 71: Reducing Variation – Multiple Readings
  If improving of gage instructions and appraiser training fail to provide an acceptable outcome you can use multiple readings to gain a better result
  One drawback to this is that it can take significantly longer to conduct the study and this method will need to be used until further gage improvements can be made

## Slide 72: Reduction Variation – Multiple Readings
  Variation may be reduced by taking multiple measurements of a part and averaging the readings.
  This average can then be substituted for the individual measurement
Where N = Sample size

## Slide 73: Multiple Readings – Why?
Measurement
Actual Size
Confidence Interval
Average Measurement

## Slide 74: Multiple Readings - Example

## Slide 75: Multiple Readings - Example
Five measurements were taken and averaged to help reduce variation
Initial GRR (1 Measurement) | 30.79 %Study
Repeat GRR (5 Measurements) | 16.82 %Study

## Slide 76: Attribute Data GRR

## Slide 77: Attribute GRR
What is the goal of an attribute GRR?
  100% match within and between operators measurement and the correct (master) attribute
  To have all Kappa values greater than 0.75

## Slide 78: Attribute GRR
What you need
  At least 50 parts
  At least 3 appraisers
  Each appraiser should measure each part 3 times

## Slide 79: Attribute GRR Part Selection
Chart: 
30% from normal process variation
25% close to the LSL
25% close to the USL
10% outside of USL
10% outside of LSL

## Slide 80: Attribute GRR
How to conduct an attribute GRR
  Select parts for study (50 minimum)
  Select appraisers for study (recommend 3)
  Measure all 50 parts on a master measurement system (ex. CMM)
  Build a master truth table
    This table lists the known disposition of each part (pass/fail) as well as associated variable data

## Slide 81: Attribute GRR
Conduct the GRR study (Measure parts in random order, observe study while in progress.
Input data into Minitab
Input part number, master measurement and truth
Input trial number, appraiser and outcome

## Slide 82: Attribute GRR
1. Go to ‘Stat’, ‘Quality Tools’, ‘Attribute Agreement Analysis’
Select the appropriate columns for each field.
Don’t forget to fill in the gage information!
If appraiser data is in multiple columns fill in here

## Slide 83: Assessment Agreement - Attribute
How well each appraiser agrees with themself
How well each appraiser agrees with the truth

## Slide 84: Attribute Agreement Analysis for Pass/Fail
Kappa Values
(Must be greater than 0.75)

## Slide 85: Between Appraisers
Kappa Values
(Must be greater than 0.75)

## Slide 86: Kappa Value
What is a kappa value?
  Fleiss’ kappa value is used to measure the level of agreement between two appraisers rating the same data set
  This equation is comparing the proportion of agreement between appraisers after removing chance

## Slide 87: Attribute Agreement
How many Kappa values are calculated?
  Within Appraiser
  Between Appraiser
  Each Appraiser to Standard
  All Appraisers to Standard
All Kappa values must be greater than 0.75
Notes: Fleiss’ Kappa value is generally accepted as being more robust than a simple percent agreement as the kappa value takes into account the agreement potentially occurring by chance.
Cohen’s kappa value is not generally used as you can only have 2 appraisers and one trial per appraiser.

## Slide 88: Attribute GRR – More Info
How many Kappa values are calculated?
  Within Appraiser
  Between Appraiser
  Each Appraiser to Standard
  All Appraisers to Standard
All Kappa values must be greater than 0.75
Notes: Within – Repeatability
Between – Reproducibility

## Slide 89: Guard Banding

## Slide 90: Guard Banding – General Rules
Guard banding is shrinking the product acceptance limits from the product tolerance
This can only be done if a gage has repeatability issues. It will do nothing to correct a reproducibility concern
Typically used when a product is not capable and the part is measured 100%

## Slide 91: Guard Banding – How to Implement
Chart: 
Let’s assume a gauge has 20% Tolerance Variation measurement error (GRR)
What would you do?
10%
10%
Both the LSL and USL should be shrunk by 10% (toward nominal)

## Slide 92: Acceptance Criteria - Summary
Resolution
  Gage should be 1/10th of product tolerance
  Range charts with 5 or less possible values (steps) could be an indicator of inadequate resolution
  Range charts where 25% or more points have 0 range values indicates poor resolution
Linearity
  Must include 5 or more samples
  Parts selected must cover entire product tolerance
  All p-values should be greater than 0.05 (P>0.05)

## Slide 93: Acceptance Criteria - Summary
Variable GRR
  10% or less %Study with 5 or more distinct categories (10%-30% may be acceptable)
  10% or less %Tolerance with 5 or more distinct categories (10%-30% may be acceptable)
  If reporting %Tolerance process must have high capability (ex Ppk>2.5), be stable and in control. Number of distinct categories is not used when reporting %Tolerance with a high capability process
  R-Charts should have all points within control limits
  Xbar charts should have more than 50% of points outside of control limits.
  Part to part variation should account for almost all of the variation in the study

## Slide 94: Acceptance Criteria - Summary
Attribute GRR
  All Kappa values must be greater than 0.75
    Within Appraiser
    Between Appraiser
    Each Appraiser to Standard
    All Appraisers to Standard